2017 Publication guidelines for structural modelling of small-angle scattering data from biomolecules in solution: an update given progress since 2012

Jill Trewhella1*, Anthony Duff2, Dominique Durand3, Frank Gabel4, J. Mitchell Guss1, Wayne Hendrickson5, Greg L. Hura6, David A. Jacques7, Nigel M. Kirby8, Ann H. Kwan1, Javier Pérez9,Lois Pollack10, Timothy M. Ryan8, Andrej Sali11, Dina Schneidman-Duhovny12, Torsten Schwede13, Dmitri Svergun14, Masaaki Sugiyama15, John A. Tainer16, Patrice Vachette17, Andrew E. Whitten2

1 School of Life and Environmental Sciences, The University of Sydney, NSW 2006, Australia

2 Australian Centre for Neutron Science, ANSTO, New Illawarra Rd, Lucas Heights, NSW 2234, Australia

3 Institut de Biologie Intégrative de la Cellule, UMR 9198, Bâtiment 430, Université Paris-Sud, F-91405 ORSAY Cedex, France

4 Institut de Biologie Structurale (IBS), Université Grenoble I, Grenoble, France, Centre National de la Recherche Scientifique (CNRS), Grenoble, France, Commissariat à l'Energie Atomique (CEA), Grenoble, France

5 Biological Sciences, Columbia University Medical Center, New York, USA

6 Physical Biosciences Division, Lawrence Berkeley National Laboratory, Berkeley, California, USA.

7 University of Technology Sydney, ithree Institute, 15 Broadway, Ultimo NSW, 2007, Australia

8 Australian Synchrotron, 800 Blackburn Rd, Clayton VIC 3168, Australia

9 Institute for Integrative Biology of the Cell (I2BC), CEA, CNRS, Universite´ Paris-Sud, Universite´ Paris-Saclay, Gif-sur-Yvette, F-91198, France

10 Applied and Engineering Physics, Cornell Universitym, 254 Clark Hall, 142 Sciences Drive, Ithaca, NY 14853-3501, USA

11 Departments of Bioengineering and Therapeutic Sciences, Pharmaceutical Chemistry, and California Institute for Quantitative Biosciences (QB3), University of California at San Francisco, San Francisco, CA, USA

12 Biological Chemistry, Hebrew University of Jerusalem, Israel

13 Biozentrum, University of Basel & SIB Swiss Institute of Bioinformatics, Basel, Switzerland

14 European Molecular Biology Laboratory (EMBL) Hamburg, c/o DESY, Nokestrasse 85, Hamburg, 22607, Germany

15 Research Reactor Institute, Kyoto University, Kumatori, Sennan-gun, Osaka 590-0494, Japan

16 Basic Science Research Division, Molecular & Cellular Oncology, MD Anderson Cancer Center, University of Texas, Houston TX, USA

17 Institut de Biologie Intégrative de la Cellule, UMR 9198, Université Paris-Sud, F-91405 ORSAY Cedex, France

*Correspondence email: jill.trewhella@sydney.edu.au

Abstract

Five years ago we published preliminary guidelines addressing sample quality, data acquisition and reduction, presentation of scattering data and validation, and modelling for biomolecular small-angle scattering (SAS) experiments. Biomolecular SAS since has continued to grow and develop and authors increasingly have adopted the preliminary guidelines. At the same time, integrative/hybrid methods for structural modelling of increasingly complex biomolecular complexes and assemblies is a rapidly growing field that is pushing at the frontier of structural biology. In order for SAS to realize its full potential contribution to this field, it is essential to ensure open access to the information required for evaluation of the quality of samples and data, as well as the validity of structural models. To this end, we review here and update the preliminary guidelines for data presentation in a publication, and recommend deposition of data and associated models to a public archive. These guidelines and recommendations have been prepared in consultation with members of the International Union of Crystallography (IUCr) Small-Angle Scattering and Journals Commissions, the world-wide Protein Data Bank (wwPDB) Small-Angle Scattering validation task force, and additional experts in the field.

Introduction

The objective of putting out the preliminary guidelines for biomolecular SAS experiments (Jacques et al., 2012a) was to provide a reporting framework so that “readers can independently assess the quality of the data and the basis for any interpretations presented.” The focus was on solution SAS experiments where the primary goal is the generation and testing of 3D models (Jacques et al., 2012b). These guidelines were developed in consultation with members of the SAS and Journals Commissions of the IUCr and other experts in the field, are now used by many authors and are endorsed by the IUCr Journals (http://journals.iucr.org/services/sas/).

Since the preliminary publications appeared, the world-wide Protein Data Bank (wwPDB) established the Small-Angle Scattering validation task force (SASvtf, https://www.wwpdb.org/task/sas) that has made recommendations regarding the archiving and validation of SAS data and models (Trewhella et al., 2013). Further, the wwPDB Integrative/Hybrid methods (IHM) validation task force has been formed (Sali et al., 2015) to address the complex issues concerning the archiving and validation of models of biomolecular complexes and assemblies that depend upon data obtained from independent experimental techniques, including SAS data, often representing information at different scales (Trewhella, 2016). There also have been substantial advances in analysis tools for SAS (Chen and Hub, 2015; Franke et al., 2015; Konarev and Svergun, 2015; Petoukhov et al., 2012; Petoukhov and Svergun, 2015; Rambo and Tainer, 2013; Schneidman-Duhovny et al., 2013) and instrumentation, in particular the growth of SAS experiments utilizing in-line purification and characterization (Blanchet et al., 2015; Bras et al., 2014; Brookes et al., 2013; Brookes et al., 2016; Graewert et al., 2015; Ibrahim et al., 2017; Jordan et al., 2016; Meisburger et al., 2016). In regard to modelling SAS data, there has been significant increased interest and methods development in multi-state/ensemble-based methods for flexible biomolecules (Berlin et al., 2013; Kikhney and Svergun, 2015; Perkins et al., 2016; Schneidman-Duhovny et al., 2016; Terakawa et al., 2014; Tria et al., 2015) and in direct co-refinement of SAS and NMR data (Schwieters and Clore, 2014). The latter places especially stringent requirements on the accuracy and precision of SAS data.

The recommendations of the SASvtf (Trewhella et al., 2013) have progressed substantially in regard to model validation and archiving. In particular, work has begun in support of community discussions and technical developments required to develop a federated system of data banks to support dissemination and validation of integrative/hybrid models (Sali et al., 2015). In particular:

- A standard dictionary with definition of terms for collecting and managing SAS data as well as facilitating data exchange between laboratories and data banks has been developed (Kachala et al., 2016), building upon the sasCIF (Malfois and Svergun, 2000) that was developed originally as an extension of the core Crystallographic Information Framework (CIF);

- A freely accessible and fully searchable SAS experimental data and model data bank (SASBDB, https://www.sasbdb.org/ (Valentini et al., 2015)) has been established to be part of an envisioned federated system of interoperable databases supporting hybrid data and model validation.

The SASvtf report reiterated the importance of the recommended preliminary publication guidelines and expanded on them, further stating that “criteria need to be agreed upon for the assessment of the quality of deposited data and the accuracy of SAS-derived models, and the extent to which a given model fits the SAS data.”

In the light of the above developments, it is timely to update the preliminary publication guidelines. We have followed the same structure as previously, with four sections covering (1) sample quality, (2) data acquisition and reduction, (3) presentation of scattering data and validation, and (4) modelling. Each section describes briefly the relevant context with a tabulated summary of the specific information to be reported. Importantly, we have added to our recommendations that SAS data and models, along with the details of the experiment as described in each of the four sections here, be deposited in a public data bank. An example report is provided at the end of a specific set of SAS experiments in a form that is consistent with the guidelines and demonstrating the value of complete reporting. While many of the recommended guidelines are best practice for biomolecular SAS generally, our main focus remains on experiments aimed at three-dimensional (3D) structural modelling from solution SAS data. As such, SAS experiments aimed at understanding highly heterogeneous mixtures or transient species using time-resolved data are not explicitly considered.

Importantly, the guidelines are not intended to restrict publication, but rather to ensure adequate description of the accuracy and confidence in the data and modelling outputs. The user of the information can then, for example, distinguish between data and modelling results that may have evident imperfections but nonetheless provide useful qualitative structural insights, and quantitative parameters that can contribute to hybrid methods modelling where an objective is to understand the accuracy of components of a model that have different kinds of contributing data (Schneidman-Duhovny et al., 2014; Yang et al., 2012).

(1) Sample Quality

Given the paramount importance of sample preparation and characterization for biomolecular structure modelling using SAS data, sample quality must continue to be emphasized. A SAS profile can be measured from any sample, and unlike crystallography and NMR where there are both quantitative standards and internal controls for assessing sample and data quality, a SAS profile by itself does not provide sufficient information for such assessment. Fundamental to a successful interpretation of a biomolecular SAS experiment in terms of structural models is that the scattering data are demonstrated to be from mono-dispersed particles in the dilute solution regime. This means that the SAS data are free of contributions from contaminants and the effects of non-specific aggregation or inter-particle distance correlations. To avoid these systematic biases, well-characterized solutions of high purity must be measured, yielding SAS profiles that encode information regarding biomolecular structure (the form factor). Additionally, as coherent scattering that encodes the desired structural information for a biomolecule in solution is inherently weak (~1 in 106 incident photons are scattered from a 1 mg/mL solution of a 15-kDa protein (Stuhrmann, 1980)), accurately and precisely scaled measurements with respect to incident radiation of solvent plus biomolecule and precisely matched solvents also are essential. As described in the following sections, an inaccurate solvent subtraction from the solvent plus biomolecule of interest will affect important validation parameters and structural interpretation.

Traditionally, solution SAS data have been collected at multiple concentrations of the particle of interest to evaluate and eliminate concentration-dependent biasing effects through strategic choice of solvent conditions or extrapolation to infinite dilution. The molecular mass (M) or volume (V) of the scattering particle then can be estimated from zero-angle scattering, I(0). Historically, researchers have used relative I(0) values compared to a reference protein to estimate M (Mylonas and Svergun, 2007). Alternatively, one can avoid uncertainties introduced by reference to another protein, by placing SAS data on an absolute scale by reference to pure H2O scattering or incident beam flux (Jacrot and Zaccai, 1981; Orthaber et al., 2000). Determination of M or V from I(0) requires accurate concentrations of the sample constituents, which can be challenging. With careful choice of solvent constituents, a well-determined UV extinction coefficient, and a properly calibrated UV spectrophotometer, the accuracy of the estimated concentration can be better than ~10%. Agreement of the I(0)-based estimate for M with that determined from the scattering particle chemical composition (including protein or nucleic acid sequence with purification tags if still present plus any cofactors, modifications or bound ligands, and in the case of small-angle neutron scattering (SANS) isotopic composition) is an important demonstration that the measured SAS profile corresponds to the form factor of the particle of interest, is free of non-specific associations and in the dilute solution approximation.

Developments of in-line purification and characterization of samples at synchrotron small-angle X-ray scattering (SAXS) (e.g. (Brennich et al., 2017; David and Perez, 2009; Graewert et al., 2015; Mathew et al., 2004)) and at SANS (Jordan et al., 2016) beam lines have led to an increasing uptake of size-exclusion chromatography (SEC) prior to injection into the SAS measurement cell or capillary to separate out the components of mixtures and poly-disperse solutions, or in the case of membrane proteins detergent micelles or nanodiscs (Berthaud et al., 2012). This combined SEC-SAS approach has been extremely successful at synchrotron SAXS facilities and opened up studies of systems that previously were impossible due to time-dependent aggregation. A drawback to the approach is the necessary dilution of the sample on the SEC column. Additionally, as the fluid in the center of tubing linking the SEC column to the SAXS cell flows faster than that at the edges of the tube (Poiseuille flow), the SEC peak will broaden before measurement, and depending on how optimized the system is the broadening can be quite significant. Depending on the set up, it can be difficult to accurately correlate UV absorbance measurements for the sample coming off the column to a given SAS measurement frame, thus limiting the ability to determine sample concentrations. Monitoring UV absorbance immediately prior to SAS measurement with minimal intervening path-length and volume, or ultimately co-incident measurement, is facilitating increased accuracy in the estimation of M or V from I(0).

Accurate solvent subtraction is essential for the complementary method for estimating M from the scattering invariant (Qi) (Porod, 1951) and its relationship to the scattering particle volume (Debye et al., 1957; Porod, 1951) or for assessing the presence of flexibility in the scattering particle (Kratky, 1982). Solvent blank mis-match with the sample will introduce errors that will confound these analyses as they depend on an accurate representation of the scattering at high angles that are orders of magnitude less intense than the lowest angle data for the protein and are only a few % above the solvent scattering. For SANS data, contributions to the background from incoherent scattering of 1H also can prove problematic as it is 10-20 times the total scattering (coherent plus incoherent contributions) from other nuclei present in a biomolecule (Jacrot, 1976) and difficult to control with sufficient precision. As a result solvent subtractions for SANS data with significant 1H often include adjustments by an arbitrary addition or subtraction of a constant to force the scattering at high angles to approximately zero. The need for this adjustment can be minimized by using a final dialysate as the solvent blank where the dialysis has been maintained in a closed environment to avoid differential 1H-2H exchange between the sample and dialysis reservoir, and by calibrating sample and solvent transmissions against pure 1H2O and 2H2O.

Excellent descriptions of preparing quality samples and well-matched solvent blanks for SAS experiments that demonstrably meet the stringent requirements for obtaining SAS data suitable for structural studies and for preparing for a synchrotron and neutron contrast variation experiments have recently appeared in Nature Protocols (Jeffries et al., 2016; Skou et al., 2014). Together these papers provide important and comprehensive practical advice for the preparation of a SAS experiment.

Reporting guidelines for demonstrating sample quality

Source of samples, including sample purification protocol, a measure of the final purity and how it was determined.

Composition of the sample, including sequences, or references to them, for the proteins and nucleic acids with purification tags if unable to be removed, any ligands, cofactors or other modifications and the predicted molecular mass. FASTA IDs with the relevant amino acid ranges specified as appropriate.

Solvent/buffer pH and composition, including additives such as free radical scavengers used to minimize the effects of radiation damage during SAXS data acquisition, and a statement of how the SAS measured solvent blank was obtained (e.g. last step dialysate, concentrator or column flow through.)

Sample concentration(s) and method(s) of determination, including extinction coefficients and wavelengths when UV absorbance measurements are used.

In the case of combined SEC-SAS experiments, description (or reference) to system, column size/type/resin, injection concentration (mg/mL) and elution volume (mL), flow rate (mL/min).

In the case of SANS contrast variation experiments, sample and buffer transmissions referenced to transmissions of pure 1H2O and 2H2O, accurate percentage of deuteration of each biomolecular component (e.g. from mass spectrometry) and of the solvent (e.g. from densitometry, weighing, transmissions).

Any SAS-independent assessments of mono-dispersity over a range of conditions (e.g. Analytical Ultracentrifugation, Dynamic Light Scattering and/or aggregate-free gel filtration and/or Multi-Angle Laser Light Scattering) that complements the SAS-based assessments.

(2) Data Acquisition and Reduction

In the case of isotropic solution scattering, data reduction refers to the process of converting counts on a 1- or 2-D detector to the scattered intensity profile arising from the sample, with associated errors, as I(q) versus q (where q=4πsinθλ, and  2 is the scattering angle,  the wavelength of the radiation). To obtain the SAS profile relating to the structure of the particle of interest, the data reduction software must take into account detector sensitivity and non-linearity, transmitted intensity, checks for radiation damage in the case of SAXS, and normalization of intensity profiles to ensure accurate and precise subtraction of solvent scattering. Dilute solution measurement places severe requirements on monitoring beam intensity, which today can be within 0.1% and fully satisfactory. All of these procedures are described in detail in Chapter 2 from (Svergun et al., 2013).

The data reduction process also deals with potential instrumental “smearing” effects on the SAS profile (see Chapter 4 in (Glatter and Kratky, 1982)). The theory guiding interpretation of SAS data in terms of structure generally assumes a point-source geometry and a single wavelength. The instrument setup used for a SAS experiment may be an excellent approximation to the point source assumption, or may differ significantly from it requiring corrections to be made to data or to model scattering profiles for direct comparison of model with experiment. The wavelength resolution (Δλ/λ) for SAXS (whether synchrotron or laboratory-based) is generally a good approximation to a single wavelength, while for SANS it can be of the order of 1015% in order to maximize the neutron flux on sample (see examples at https://www.ill.eu/instruments-support/instruments-groups/groups/lss/more/world-directory-of-sans-instruments/). Beam size and shape also play a key role in smearing. Modern synchrotron beams and most laboratory based instruments have sufficiently small beam dimensions (in the range of tenths of a millimeter to millimeters at the detector) such that smearing effects can be safely ignored for most applications, while neutron beam dimensions can be as large as 100 mm at the detector and thus cause significant instrumental smearing.  Some laboratory-based SAXS instruments use line focused sources to maximize X-ray flux on sample. These types of instruments, first implemented by Otto Kratky (see Chapter 3 of (Glatter and Kratky, 1982)), have since been further developed for laboratory-based SAS applications (e.g. (Bergmann et al., 2000)) and data treatments must deal with significant instrumental smearing effects.

During data reduction, the SAS intensity data also can be placed on an absolute scale in units of cm-1 by comparison with the incident beam flux or scattering from pure H2O (Jacrot and Zaccai, 1981; Orthaber et al., 2000). Pure H2O is a readily accessible, universal standard whose scattering is well-characterized over a wide range of temperatures. Secondary standards are also available, such as glassy carbon (see the new NIST Standard Reference Material 3600 https://www-s.nist.gov/srmors/view_detail.cfm?srm=3600 (Allen et al., 2017).  Absolute scaling enables direct comparison of SAS data from different instruments, including X-ray and neutron sources, without arbitrary scaling and also enables determination of M or V from I(0) without reference to the scattering from a reference protein. In the case of SANS, it has been routine to place the data on an absolute scale. The practice for SAXS experimenters has been more commonly to provide data on an arbitrary relative scale, which we discourage for reasons that will be addressed further below.

Reporting guidelines for data acquisition and reduction

Instrument type (e.g. manufacturer and model designation or beamline) specifying the source (sealed tube, rotating anode, metal jet, synchrotron, spallation neutron source or reactor) and configuration used (point or line source, collimation details, detector details). In the case of SANS there may be several configurations (e.g. multiple detector positions, number of guides, apertures etc.) for a single experiment.

Beam dimensions and wavelength resolution (Δλ/λ) with data smearing parameters and measured q-range.

References to documentation for detector characteristics, the basis for error estimates, e.g. Poisson counting statistics or not, methods for detector sensitivity and linearity corrections.

Number of sample exposures and exposure times, with the method for normalization of measurements to constant counts on sample and monitoring radiation damage in the case of SAXS.

Details of the sample environment, including measurement temperature, measurement cell type and path-lengths, loading/exposure details (volumes, flow rates), any special parameters controlled, e.g. pressure, and additional in-line characterization capabilities as appropriate.

Standards measured and controls and method for placing SAS data on an absolute scale in cm-1, e.g. by reference to a well-characterized standard such as H2O or the incident beam flux. As appropriate, any standard protein measurement used as an overall check of the experimental set up.

Data reduction protocol and software used, including version number.

(3) Data Presentation, Analysis and Validation

In order for a reader to be able to assess the quality of SAS data and its suitability for structural modelling, it is necessary that the data be presented in a clear, well described manner along with the parameters and analyses that support the conclusion that the SAS profile represents the shape of the particle of interest, or in the case of flexible systems, the population weighted average SAS profile for the ensemble of conformations present.

Because I(q) decreases by several orders of magnitude over the measured q-range, data should be presented as LogI(q) vs q and/or LogI(q) vs logq. The former provides a clear representation of the data over the entire q-range, while the latter will have a zero slope at low-q if the minimum measured q-value meets the requirement of being sufficiently small to ensure adequate characterization of the largest particles present. A linear Guinier plot (Guinier, 1939)(lnI(q) vs q2) is a necessary but not sufficient demonstration that a solution contains mono-dispersed particles. The q-range upper limit for the linear Guinier approximation varies depending on particle shape and homogeneity. For a sphere of uniform scattering density, Guinier showed the limit was qRg<1.3, while for extended shapes and/or inhomogeneous particles, this limit can be < 1.0  (see Chapter 3, pp 68-73 (Fiegen and Svergun, 1987)). Assessment of the appropriate Guinier limit will be aided by complementary analyses for particle shape, such as P(r) (see below). The lower q-limit for the Guinier analysis should be the lowest, reliably measured q-value. For a particle with maximum dimensions dmax, the minimum q value measured must be at least   ~π2dmax  for accurate assessment of the particle size and shape, and as a general principle it is essential to measure below this limit to have assurance that there are no larger particles present. It has been common practice to truncate data at low-q when there are small amounts of large M impurities, aggregation or polydispersity present, resulting in some upturn of the Guinier plot. This practice is not encouraged, but in the event that it is done, it must be reported and justified. Truncating the most obviously affected lower-q data in the Guinier plot will not completely eliminate the effects of the contaminant and will thus have an effect on the derived structural parameters that must be taken into account. For Guinier fits, a quality-of-fit parameter such as the Pearson’s residual (R) or coefficient of correlation (R2) value for a linear fit is the most useful to report.

The Fourier transform of the scattering profile yields P(r) vs r, the scattering density difference-weighted distribution of distances r between atoms and is generally computed as the indirect Fourier transform of I(q) (Glatter, 1977). By its definition, P(r) is equal to zero for r values exceeding the maximum particle size dmax. Agreement between the P(r) and Guinier-determined Rg and I(0) values is a good measure of the self-consistency of the SAS profile as P(r) is calculated using a larger portion of the measurement range. That said, it is not correct to simply choose a dmax that provides a solution that agrees with the Guinier Rg. Rather the P(r) solution must be independently optimised with the understanding that dmax is an input parameter to the indirect transform selected by the user based on the observed fit of the model I(q) corresponding to a given P(r), how P(r) approaches zero at r = 0 and dmax and what might be known about the dimensions of the system from complementary techniques. The indirect Fourier transform methods for calculating P(r) include a smoothing parameter that is a complicating factor in assessing the quality of the fit for a given solution. A simple χ2 test is straightforward to calculate, though it does have limitations as will be discussed in the 3D Modelling section below. One approach used by the popular program GNOM for calculating P(r) is to use a quality-of-fit assessment (referred to as the “total estimate”) that is based on χ2 combined with a number of “perceptual criteria” (Svergun, 1992).

The molecular mass, M, in Daltons for a scattering particle is readily calculated as:

M=I0NAC(ΔρM)2 							(1)

where I(0) is on an absolute scale in units of cm-1, NA is Avogadro’s number, C is the concentration of the scattering particle in g/mL, and ΔρM is the scattering mass contrast ∆ρϑ where Δρ is the average scattering length density difference between the particle and its solvent in cm-2 (or cm/cm3, scattering length/unit volume), and ϑ its partial specific volume in cm3/g (Orthaber et al., 2000). Δρ and ϑ are both related to the molecular volume and can be readily estimated in a consistent manner for both X-rays and neutrons from the chemical and isotopic composition of the particle and its solvent in a standard way. There are several web-based tools for calculating Δρ and ϑ (see Contrast model of MULCh (http://smb-research.smb.usyd.edu.au/NCVWeb/index.jsp) for Δρ and ϑ for SAX and SANS, also the web version of US-SOMO https://somo.chem.utk.edu/somo/ will calculate ϑ and other molecular properties rapidly from the sequence).

https://somo.chem.utk.edu/somo/ For an easy to use protocol for the M calculation see Box 2 in Jeffries and colleagues (Jeffries et al., 2016).

In the case of SANS, the relationship between I(0) and molecular volume is often formulated as:

I0=Cbt-ρsolV2						(2)

Where bt is the sum of all scattering lengths in the scattering particle, ρsol is the scattering length density of the solvent and V the particle volume. Here bt and ρsol are readily calculated from chemical composition, and thus one can extract V using the measured values for I(0) and C and only need to estimate the partial specific volume if you want to estimate M from V.

Historically, proteins have been used as a calibration standard for estimating M. From equation (1) it is seen that if the product of Δρ and  is assumed to be the same for all proteins, mass is proportional to I(0) normalized by protein concentration in w/v units (Mylonas and Svergun, 2007). However, the simplest implementation of this ratio method is not readily applicable to polynucleotides. For proteins experimentally determined values for ϑ vary by as much as 10% and Δρ also can be quite variable, especially in the case of bound metal ligands for example. Additionally, it is the case that most readily available inexpensive protein standards have some tendency for time- and/or radiation-induced aggregation or degradation which introduces further systematic error in the assessed M value. It nevertheless can be useful in practice to measure a known protein standard (such as lysozyme, bovine serum albumin or glucose isomerase) as a check of the overall experimental set up. However, dependence on this approach for the evaluation of M is discouraged in favor of absolute scaling of the SAS data and either Eq. (1) or (2) above.

The total scattered intensity (calculated as the integral from zero to infinity of q2I(q)) is referred to as the Porod invariant Qi  that, for uniform scattering density particles with a well-defined boundary, depends only on  the volume of the scattering particle and not its form (Porod, 1951).,The particle volume or Porod volume, VP, then is calculated as:

Vp=2π2I(0)/Qi							(3)

As Qi is an integral from zero to infinity it cannot be determined directly from the measured scattering profile without extrapolating to zero in one direction and infinity in the other. In practice, the scattering from a uniform scattering density particle is zero for q values where information on the particle shape has ended, which typically is somewhere in the range 0.10.4 Å-1 and the integral is generally estimated using a smoothed, regularized scattering profile obtained from P(r) (e.g. from GNOM (Svergun, 1992)) that includes a correction (in the form of a solvent subtraction) to force the high q data obey the expected q-4 dependence for a uniform scattering density particle with a well-defined boundary (Porod, 1951). By interrogating a large set of theoretical scattering profiles calculated from coordinates of proteins in the Protein Data Bank (PDB) (Berman et al., 2000), Fischer and colleagues (Fischer et al., 2009) determined empirical correction factors for estimating Qi for scattering data acquired over specific measured q-ranges. Rambo and Tainer defined a volume of correlation, Vc, method that is applicable to unfolded or flexible proteins (Rambo and Tainer, 2013). Any one of these methods can be used to estimate the volume of the scattering particle, which can then be related to M, keeping in mind that they all are highly dependent on an accurate background subtraction.

For a typical folded and hydrated protein, a useful rule of thumb is that M/VP is ~1.45 - 1.5. Agreement of this estimate with that derived using Eq. (1) and with the expected value from the chemical composition of the particle of interest (full sequences, including tags, bound ligands and modifications) is a primary validation parameter that demonstrates the scattering particle is a mono-disperse, folded macromolecule or macromolecular complex, and that the SAS data are suitable for structural interpretation in terms of a 3D model.

In the case of SANS with contrast variation data, I(0) and Rg values vary with contrast and should be reported for each contrast point measured. The M or V estimate from I(0) also should be determined for each contrast point to identify potential 1H2O induced aggregation effects (from eq. (1), for a constant M and , I(0) 1/Δ2). In addition, the Stuhrmann plot (Rg2 vs 1/Δ (Koch and Stuhrmann, 1979) is valuable to show as it provides information on internal scattering density variations within the scattering particle. For a particle composed of discrete components with distinct mean scattering densities (e.g. a protein plus polynucleotide, or 2H-labelled and unlabeled proteins) a combination of the Stuhrmann analysis and application of the parallel axis theorem will provide information on the disposition of components, their Rg values and the Rg value for the total particle at infinite contrast where internal scattering density fluctuations are negligible. In addition to these analyses, the scattering functions encoding information on the individual components (from the scattering from pairs of scattering centers within each component) and their dispositions (from the scattering from pairs of scattering centers scattering from different components, or the cross term)  can be extracted for standard Guinier and P(r) analyses and modelling. The MULCh suite of programs (ModULes of the analysis of Contrast variation data, available for download and as a web-based tool at http://smb-research.smb.usyd.edu.au/NCVWeb/index.jsp (Whitten et al., 2008)) was designed to aid in planning a contrast variation experiment by providing the dependence of I(0) on contrast for given deuteration levels in biomolecular components and solvent (Contrast module), for Sturhmann and parallel axis theorem analysis (Rg module), and for extraction of component scattering function, including the cross term (Compost module).

The above q-4 approximation for the decay of high-q data is a good approximation for most folded proteins, but not for unfolded proteins where the dependence is q-2 (Debye, 1947).  The asymptotic behavior of the high-q data thus can distinguish between folded, partly flexible and unfolded structures. Where flexibility is a possibility, its qualitative evaluation can be made using Kratky (q2I(q) vs q, see Chapter 11 (Glatter and Kratky, 1982)) and Porod-Debye (q4I(q) vs q4) (Debye et al., 1957) plots of the data (recently reviewed in (Rambo and Tainer, 2011) provided background subtractions are accurate. The dimensionless Kratky plot is extremely useful to distinguish between different degrees of folding. Proteins containing folded domains display a bell shaped curve, with a maximum of about 1.1 at around qRg = 1.75. With increasing elongation and degree of unfolding, the curve maximum shifts to the upper right and the slope of the right side of the curve increases.

It should be noted here that power laws to describe the high-q scattering that might be applicable in SAXS may in practice not be so readily applied to SANS data in 1H2O due to the high incoherent scattering signal.

Data presentation and reporting guidelines for analysis and validation

Difference scattering profiles ([particle+solvent] – [solvent scattering]) corresponding to the particle form factor should be deposited in a publically available archive or made available as supplemental material and presented in a publication as a plot logI(q) vs q or logI(q) vs logq along with a Guinier plot and: - intensities on an absolute scale in units of cm-1 normalized to unit concentration with propagated standard errors () (note: for Guinier plots (ln(I(q)) vs q2) the error in lnI(q) is)σI(q)I(q) - for multiple curves on the same plot, data can be offset for clarity with the offsets specified in the figure caption, - for SANS contrast variation experiments, data from all contrast points should be made available, - Guinier-fits displayed in the range from a minimum q~π2dmax to maximum q such that qRg < 1.3 for a globular particle, or for sufficiently asymmetric particles qRg < ~1.0, along with the quality-of-fit parameter such as Pearson’s R factor or coefficient of correlation (R2), and the associated Rg and I(0) values with errors, and - should any data from the measurement range be truncated this should be indicated along with the justification.

P(r) versus r with associated Rg and I(0) (with errors) and dmax values is essential for SAXS data and advised for SANS data (especially at solvent match points for isotopically labelled complexes) to assess data self-consistency across the measurement range. It is essential if used for bead modelling, along with the quality-of-fit parameter and its associated reference describing how it is evaluated.

M or V estimates based on I(0) plus VP, as well as values and uncertainties in the experimental parameters used in the determination and calculated parameters, such as partial specific volumes, scattering density contrasts based upon chemical composition, along with the methods of calculation or measurement.

Where applied, the magnitude of corrections for solvent subtraction, relative to the solvent scattering, applied to the data as a potential warning that something is not correct if unduly large (say 10% of the solvent scattering level).

If the presented data are desmeared (to correct for beam geometry and/or polychromaticity), this fact must be reported along with the method for desmearing and the original smeared data be made available.

For SAS data measured on solutions at multiple concentrations of the biomolecule of interest to evaluate concentration-dependent behavior indicative of aggregation or inter-particle interference; if no change in Rg or I(0)/C is observed with increasing concentration (C, in units of w/v), this fact should be noted. Best practice would be to report M estimates at different concentrations as evidence that the analyzed data are in the dilute approximation range or to provide a plot of I(0)/C vs C.

A dimensionless Kratky plot as a check on the degree of folding and/or flexibility in the scattering particle. Kratky and/or Porod-Debye plots might also be used to assess potential flexibility.

For SEC-SAS data, in the case of single well-resolved peaks, a plot of I(0) and Rg as a function of measurement time or measurement frame, and correlated UV traces if used for estimating concentrations. For more complex cases, deconvolution may be needed, e.g. using the HPLC-SAXS module of US-SUMO http://www.somo.uthscsa.edu/.

Data processing used to obtain the final data set for analysis and modelling (including data reduction to I(q) vs q, solvent subtraction, merging of multiple data sets, extrapolation to infinite dilution, etc). For such processed data, the original measurements should be deposited in a publicly available data bank along with the precise protocol used for processing.

For contrast variation experiments the nature and number of contrast points with a plot of normalized ±√I(0) versus solvent density identifying the total particle solvent match point along with transmissions at each contrast with controls for pure 1H2O and 2H2O for calibration.

For the contrast variation experiments of assemblies of components with different mean Δ, M or V estimates from I(0) for each contrast point, Stuhrmann plots and derived Rg values for individual components and whole particle at infinite contrast and extracted component scattering functions (including cross-term) is desirable.

Software used for data processing and analysis, such as calculation of Rg, VP, and P(r), including version numbers.

(4) 3D Modelling

Having obtained an accurate and sufficiently precise I(q) versus q plot for the particle of interest, demonstrated it has the expected M or V value, and assessed its potential flexibility, a 3D modelling strategy can be selected. This strategy may include ab initio shape or bead modelling and/or atomistic modelling using domains or subunits of known structure, usually derived from crystallography or NMR experiments and combined with additional experimental restraints, to obtain a model that is optimized to fit the scattering data and minimizes 2.

As solution scattering data are essentially 1D questions regarding the resolution, ambiguity and precision of the resulting 3D models must be addressed. AMBIMETER is a recently released program that provides an a priori assessment of the ambiguity associated with spherically averaged single-particle scattering (Petoukhov and Svergun, 2015). In the case of data that can be adequately fit to a single average 3D model (either shape or atomistic representations) an evaluation of the inherent ambiguity in the modelling solution is required. Here a question to answer is whether a single best-fit model or class of models uniquely fit the data, or if multiple classes of models exist that fit the data equally well. This evaluation is frequently done by running multiple independent model optimizations and using a cluster analysis to compare models in terms of shapes or in the case of atomistic models, relative positions and orientations of domains or subunits and contact between the different components. The number of clusters that fit the data provides an estimate of the ambiguity, while precision is defined by the variability within the cluster. Reliable restraints from complementary experiments (e.g. domain structures from NMR or crystallography, information on symmetry, contact information and distance or orientiational restraints from NMR, cross-linking, Förster resonance energy transfer) or from bioinformatics analysis (e.g. sequence conservation or co-variation) can be imposed as part of any modelling strategy to reduce the inherent SAS data ambiguity and also the number of degrees of freedom in the optimization.

Symmetry assumptions in bead or shape modelling can highly influence the resulting models and thus if symmetry is imposed to generate a model that is to be used, it is important to demonstrate that a similar result is obtained in the absence of symmetry restraints.

If monodispersity in solution cannot be achieved or guaranteed, the measured scattering intensity reflects the spherical average over all K species present.  Assuming no interaction between the particles, the scattering intensity is then a linear combination of the scattering of the species Ik(q) multiplied by their respective non-negative volume fraction vk:

(4)

Depending on the number of components in solution, there are various approaches to data analysis. In the case of mixtures with not too many components whose individual scattering intensities are known, the volume fractions may be estimated from Eq. 4 as in the program OLIGOMER (Konarev et al., 2003). For systems with unknown structure existing in a stable equilibrium, e.g. monomer and dimer with known association and disassociation constants, 3D structural analysis is possible. This can be done ab initio or using rigid body modelling, e.g. with GASBORMX or SASREFMX (Petoukhov et al., 2013). The reporting guidelines for using these programs are similar to the monodisperse case but with the extra parameter of volume fractions of all species in solution (typically, multiple curves are recorded for analysis, e.g. a concentration series).

Perhaps the most complicated mixtures are flexible systems containing multiple conformers, e.g. multi-domain proteins with flexible linkers or hinges. For such systems, the number of terms in Eq. 4 is astronomically high. Such systems may still be characterized with multi-state or ensemble methods, where a large population of potential conformations is generated and sub-states or sub-ensembles that describe the observed scattering data based on a priori information are selected (Berlin et al., 2013; Bernado et al., 2007; Kikhney and Svergun, 2015; Pelikan et al., 2009; Perkins et al., 2016; Schneidman-Duhovny et al., 2016; Terakawa et al., 2014; Tria et al., 2015; Yang et al., 2010). As the number of degrees of freedom in ensemble modelling is so much larger than when optimizing a single average model, the danger of overfitting and over interpretation is significantly amplified. Satisfactory solution of the problem of multi-state/ensemble modelling thus depends greatly on the application of restraints from complementary experiments or bioinformatics to limit the conformational space that must be sampled. While many programs for multi-state/ensemble modelling produce representative structures to describe the range of states within the population, these representative structures are generally neither accurate nor precise and only aid in providing a visual, qualitative description of the nature of representative states. On the other hand, the distribution of Rg values for the optimized ensemble is generally quite robust, providing a quantitative measure of the extent of structural flexibility (Bernado et al., 2008; Carter et al., 2015). In cases where the conformational space is sufficiently constrained, it may be practical to evaluate the ambiguity and precision of the multi-state/ensemble models. For example, consider a system where the data are explained by “open” and “closed” structural states. A cluster analysis on the opened and closed states may reveal little variability in the closed state, thus low ambiguity and higher precision, while the open structure may show larger variation and consequently high ambiguity and low precision (e.g. see Figure 3J (Carter et al., 2015)).

For atomistic representations, the protocol used to include contributions to the scattering data from the hydration layer is important. These effects are quite significant for SAXS and for SANS from samples with high levels of D2O (Kim and Gabel, 2015; Perkins, 1986; Svergun et al., 1998; Zhang et al., 2012). They become especially significant and important to report in the co-refinement of SAXS/NMR data for solution structure determination (Grishaev et al., 2010).

The most commonly used parameter for evaluating the discrepancy between the fit of a 3D model to a scattering profile is the global fit parameter χ2 that is defined most simply as:

χ2=1N-1j=1N(Iexpqj-cImod(qj)σ(qj))2					(5)

where N is the number of points in the scattering profile, Iexpq is the experimental scattering profile, Imodq is the computed scattering profile based on the 3D model, c is a multiplicative scaling parameter that is used to minimize χ2, and σq is the standard deviation for each measured data point. From Eq. (5) we see that χ2 will be smaller for data with large errors and thus is not so much a measure of the adequacy of the model with respect to the data, but rather measure of how much can the model still be improved, given the statistical quality of the data.

Scattering data are acquired as the sum of events on a detector. A model that fits the data will have a χ2 value close to 1 providing the statistical errors are propagated correctly and there are no systematic errors. For synchrotron SAXS data, the increasing brightness of the sources has reduced the magnitude of statistical errors in the data to the extent that they may no longer dominate and systematic errors can become significant. A recent model has been proposed and tested for optimizing experimental setups and taking into account not just statistical errors, but those originating from setup geometry and the physics of the measurement process (Sedlak et al., 2017). Accurate error propagation is an important area for further work in the community for SAS data to contribute to integrative/hybrid modelling.

Over- or under-estimation of the statistical errors and potential contributions from systematic errors have led to reported χ2 values ranging from a few tenths to quite large values (>5) and yet the fits to the data may be good, even excellent, or claimed to be good based on a “by eye” evaluation of a presented plot (e.g. see supplementary Fig. 2 in (Appolaire et al., 2014). Generally, SAS intensity decreases rapidly and by orders of magnitude over the measured q-range, and depending upon how the data are presented, regions of significant mis-fit of the scattering profile may not be apparent. Also, as χ2 is only a global fit parameter, it is important to present the data and model fit so that systematic deviations that may be present in specific q-regimes are evident; e.g. in the mid-q regime most highly influenced by domain positioning and orientation. A straightforward and intuitive approach to demonstrating the quality of a model fit over the entire measured or modelled q-range of a SAS profile that takes into account relative errors is a difference plot of Iexpq-cImod(q)σ(q) vs q, as is nicely demonstrated in Figs 3, 4, and 5 from (Carter et al., 2015)). The error weighting of this difference plot aids in visualization by preventing the plot from being dominated by regions of weaker scattering and poor statistics. This plot presents the fit in the noisy high-q regions without losing information in the low to mid-q regions that contain the essential shape information that is most important for biomolecular SAS-modelling. If the deviations from the model are only evident in the high q-regime, it might be indicative of an error in the solvent subtraction or unaccounted for disorder.

Different modelling programs use various adjustable parameters in their procedures to minimize χ2 and these should be reported (e.g. for CRYSOL the parameters Dro and Ra specifying the scattering density contrast and thickness in the hydration layer model and a constant term to account for possible errors in the solvent subtraction step by forcing a q-4 dependency at high q, or for FoXS the parameters c1 and c2 are used to adjust excluded volume and hydration layer density to account for the hydration layer). Understanding these parameters is necessary to ensure that they are represent realistic assumptions given the physics of the system. Here it should be noted that not only does different modelling software use different adjustable parameters, programs sometime evolve in ways that can affect the absolute value of χ2, e.g. a later version may incorporate a constant subtraction for optimization.

The different detector characteristics, protocols for error propagation, the details of the modelling algorithm and nature of adjustable parameters renders comparisons of published χ2 values from different experiments and different modelling calculations performed at different points in time essentially meaningless.

There have been attempts made to improve on χ2 or to provide alternatives. Rambo and Tainer proposed the use of a resampling-based adaptation of the reduced χ2 test and defined a χfree2 with the aim of reducing the chance of model misidentification in noisy data and avoiding over-fitting of scattering data (Rambo and Tainer, 2013). The χfree2 parameter, however, does not solve problems relating to inaccurate error propagation.

An alternative to the traditional errors weighted residual difference of Eq. (5) is to scale differences to intensity and express the percentage difference of intensity at the corresponding q-value:

χ2=1000N-1j=1NIexp(qj)-cI(qj)Imod(qj)2					(6)

This intensity scaled χ2 can be accompanied by a plot the percentage difference in intensity (Iexpqj-cIqj*100Imod(qj) vs q providing complementary information on the relevance of the model. This plot has been implemented in the HPLC module of US-SOMO (Brookes et al., 2016).

A recently proposed alternative to χ2 that is independent of the amplitude of the statistical errors considers only the statistical likelihood of a run of consecutive points lying systematically above or below the plot generated from the fitted model (Franke et al., 2015)(program library available for download at https://github.com/GarmanGroup/CorMapAnalysis). The method has proven very useful for comparing synchrotron SAXS data frames to detect subtle radiation damage. As implemented in ATSAS, a 2D correlation map (CorMap) is generated that usefully highlights patterns of deviation. A score (P-value) is assigned relating to the statistical probability that the longest run of points that lie consistently above or below the model. While CorMap does not require knowledge of errors, because the model curve is very smooth if the errors are very small a constant sign of difference can easily be observed over a long stretch.

The above issues and limitations noted, χ2  nonetheless remains an accepted and necessary parameter to report as most modelling protocols minimize χ2 one way or another. However, reporting of a combination of χ2 values with comments on the confidence level that a global minimum was identified along with a clear representation of deviations between the model and the experimental data is essential.

Recently, an approach has been proposed to assess the resolution of ab initio models reconstructed from SAS data (Tuukkanen et al., 2016) and the method (DAMRES) is implemented in the bead modelling tools of the ATSAS package (Petoukhov et al., 2012). The method uses the Fourier shell correlation criterion widely employed in electron microscopy model assessment to estimate the resolution of an ab initio SAS models from the variability of individual reconstructions.

For a given optimized atomistic model, accuracy and precision will vary substantially for different regions depending on the contributing data. For example, the linker sequences between structured domains from crystallography or NMR that are modelled only to optimize the fit to the SAS data will not be accurate at the level of coordinate positions. Likewise interfaces not defined experimentally by crystallography or NMR will likely not be accurate. The disposition of the domains may be relatively well defined depending on how asymmetric they are and whether there were additional contacts from experiment or bioinformatics analysis that could be used as restraints, but again they will not be precise (Gabel, 2012; Kim and Gabel, 2015). As a result it is important in reporting an atomistic model to clearly identify the sources for the components of the model; where there is high resolution information and basis for building regions of unknown structure.

Reporting guidelines for modelling

All software, including version numbers, used for modelling; 3D shape, bead and atomistic modelling.

All modelling assumptions clearly stated, including adjustable parameters. In the case of imposed symmetry, especially in the case of shape models, comparison with results obtained in the absence of symmetry restraints.

For atomistic modelling, a description of how the starting models were obtained (e.g. crystal structure of a domain, homology model etc.), connectivity or distance restraints used and the basis for their selection.

Any additional experimental or bioinformatics-based evidence supporting modelling assumptions and therefore enabling modelling restraints or independent model validation.

For 3D models, values for adjustable parameters, constant adjustments to intensity as a percentage of background, χ2 and associated P-values and a clear representation of the model fit to the experimental I(q) versus q such as a residual plot or correlation map that clearly identifies systematic deviations.

Analysis of the ambiguity and precision of models based on cluster analysis of results from multiple independent optimizations of the model against the SAS profile or profiles, with examples of distinct clusters in addition to any final averaged model. In the case of ab initio bead models, an assessment of the resolution of the model reconstructions.

An example: SEC-SAXS experiments on three proteins

Figs 1-3 and Table 1 (based on the template of Supplemental Table S1) summarize the results of a set of SEC-SAXS experiments on solutions of glucose isomerase (GI, a well-characterized tetramer in solution (Ramagopal et al., 2003), bovine serum albumin (BSA, a two-domain protein with a flexible loop connecting its domains and known to be prone to oligomerization), and Ca2+-bound calmodulin (CaM, a two-domain protein known to have an extended helix with a flexible region linking the two domains (Babu et al., 1988; Barbato et al., 1992; Heidorn and Trewhella, 1988)). The examples demonstrate how the reporting guidelines aid in assembling a comprehensive description of the experiment and the models the data support. The tabulated results for all three proteins provided the input required for deposition of meta-data, data and models in the SASBDB (deposition IDs provided in Table1).

The SAXS data were acquired using the SAXS/WAXS beamline at the Australian Synchrotron (http://www.synchrotron.org.au/aussyncbeamlines/saxswaxs/saxs-specifications) with a sheath-flow sample environment to maximize X-ray dose on sample with minimal radiation loss (Kirby et al., 2016). In addition, the path-length between UV absorption and SAXS measurements was minimized, enabling the use of A280 measurements to estimate protein concentration for the SAXS data in measurement frames used for analysis (Ryan et al., in preparation). Data reduction used the beam line software ScatterBrain, while solvent subtraction, Rg, P(r), bead modelling performed using programs from the ATSAS package (Petoukhov et al., 2012), and FoXS and Multi-FoXS were used for atomistic modelling (Schneidman-Duhovny et al., 2016).

Table 1 provides the details regarding the SAXS data acquisition, samples, solvents/buffers, analysis and results. The Rg and I(0) traces (Fig. 1A) as a function of time indicate the GI and CaM samples are highly pure, as expected from their sources. High purity GI was sourced from Hampton Research and carefully stored and CaM prepared by bacterial expression and high resolution SEC (Michie et al., 2016).  In contrast, the commercially purified BSA as loaded onto the SEC column appears highly heterogeneous, consistent with the known tendency of this protein to self-associate and lack of any steps to remove higher order oligomers prior to loading.

The A280 measurements associated with the selected SAS measurement frames (Fig. S1A) for analysis were used with calculated extinction coefficients (using ProtParam, (Gasteiger et al., 2005)) to estimate protein concentrations.

Based on the Rg and I(0) traces (auto-generated by PRIMUS (Petoukhov et al., 2012)), 12 measurement frames centred on the maximum in I(0) where the Rg plot was perfectly flat were chosen for CaM data analysis. For GI, observation of Rg behaviour across the peak values in I(0) showed a small decrease in Rg after the peak (by an average of 0.6 Å over 9 measurement frames) compared to the data acquired at about half the concentration in the peak. In addition, the P(r) transform that included data from these peak frames showed the characteristic negative dip around dmax suggesting that for the buffer conditions used, even at protein concentration 0.36 mg/mL, there was small but measurable inter-particle interference contribution to the scattering. By selecting 11 frames to the right of the peak, the P(r) transform showed a much less significant negative dip. It is noteworthy that both CaM and GI are expected to have a net negative charge at pH 7.5, but only GI showed this evidence for charge repulsion. For BSA, 10 frames starting from the maximum recorded I(0) were chosen for analysis where the Rg had plateaued.

A total of 50 frames taken prior to the protein peaks were averaged for the solvent blank, although in the case of BSA this choice resulted in a slight upturn in the Guinier plot for the lowest 5 data points (q < 0.01 Å-1), which could arise either from a slight error in the solvent subtraction or aggregation. Exploration of the measurements of solvent before and after the BSA elution peak indicated variation in the solvent scattering and for BSA only, the solvent blank was taken from 50 frames after the protein had eluted. With this solvent measurement the Guinier plot was perfectly linear to the lowest measured q-value.

The logI(q) vs q plot (Fig. 1B) represents the primary SAS data, with Guinier plots shown as insets. The maximum dimensions for all the three proteins are <100 Å, and the minimum q measured (0.007 Å-1) is well below the minimum requirement of q ~π2dmax=0.015 Å-1 for accurate assessment of particles of this size. Importantly, for all three proteins, Guinier plots are linear to the first measured q-values (Pearson R-values 0.999) well below the minimum requirement for the expected dmax values and a plot of logI(q) vs logq (Fig. S1B) shows the expected zero slope at low-q expected for monodisperse particles of similar size providing confidence that the data are free of significant amounts of contaminating species and inter-particle correlations contributing a structure factor term to the scattering.

Dimensionless Kratky plots (Fig. 1C) demonstrate that the SAS data are from predominantly folded particles. GI and BSA display the expected bell-shaped curve, with a maximum of about 1.1 at around q.Rg = 1.75. The peak for BSA is slightly shifted to the right as expected for its slightly elongated shape, and the slight rise evident at qRg > 7 suggests some flexibility. The more elongated dumbbell-shaped CaM connected by an extended flexible helix gives rise to a distinct shape. The maximum on the vertical axis for CaM is somewhat higher than the expected 1.1, and shifted to qRg = 2 because of its elongated global shape while the shallow oscillation 2.5 < qRg <3.5 reflects the well-resolved two domain structure. Significant flexibility is indicated by the increase of intensity at large qRg-value (qRg = ~6). For comparison, Fig. S1 shows the standard Kratky plot.

The P(r) vs r profiles for each of the proteins (Fig. 1D) are well-behaved, showing the concave approach to zero at r = 0 and dmax expected for a mostly folded, mono-disperse protein. P(r) profiles also have the expected characteristics based on the available crystal structures; a single major peak for the globular GI and BSA, structures and the peak and shoulder expected for CaM’s dumbbell-shape. We note, that the constant adjustment applied to the background for CaM in the P(r) calculations is large at ~30% of the solvent scattering, likely reflective of the presence of significant flexibility. In contrast, GI has the smallest correction at ~3% and the value for BSA is ~8%.

For all three proteins, the Rg and I(0)-based M values (using Eq. 1) are in excellent agreement between independent Guinier and P(r) analyses (Table 1). For GI and BSA, the M values estimated from I(0) are within the expected 10% of that predicted from the sequence, whereas the CaM values are 29% larger than expected, which falls outside the expected error of ~10% based on the concentration determination. However, VP values yield M estimates within the expected uncertainty for M for all three proteins. The question that these observations pose is whether the low extinction coefficient for CaM (0.178 for a 1 mg/mL solution) combined with the shortened path-length (3.1 mm) for measurement of A280 in the SEC-SAXS line increases the error in the concentration determination.

Rg values for GI and CaM (Table 1) are in good agreement with previously published values from SAXS measurements (Guinier Rg values of 32.5 ± 0.7 Å (GI) (Mylonas and Svergun, 2007) and 21.0 ± 0.6 Å for (CaM) (Heidorn and Trewhella, 1988), whereas the value for BSA lies in between a previously published value from SAXS (29.9 ± 0.8 Å (Mylonas and Svergun, 2007)) and that predicted by the crystal structure (26.8 Å, Table 1) from the individual monomer chain A from the dimeric crystal structure (PDB:4F5S).

For all three proteins, the ab initio bead modelling program DAMMIN (Svergun, 1999) was better able to fit the data than its speedier cousin DAMMIF . However, the latter program provides a rapid assessment of the variability of the shapes that fit the data from 20 independent calculations in the form the normalized spatial discrepancy (NSD) value. The NSD value is  0.7 for GI indicating largely similar shapes, but > 0.7 for BSA and CaM suggestive of distinct classes of shape and a cluster analysis did identify 4 and 6 sub-classes for BSA and CaM, respectively. The relatively high 2 values for the DAMMIF models for GI are largely due to mis-fit around the local minimum in that profile just above q = 0.1 Å-1. Again, the BSA and CaM data have larger adjustments applied to force the high-q dependence to q-4 in the bead modelling.

As there are crystal structures for all three proteins, rather than any further assessment of the bead models, atomistic modelling was undertaken. A tetramer based on the crystal structure of GI (PDB: 1OAD) (Ramagopal et al., 2003) predicts an I(q) profile that is a reasonable fit to the scattering data (2 = 1.02) (Fig. 2). In contrast, the crystal structures of BSA (PDB:4F5S Chain A) and of CaM (PDB:1CLL) show poor fits to their respective data sets (2 = 4.4 and 10.8 respectively). There are a few missing amino acids in the CaM crystal structure (Ala-Asp-Gln at the N-terminus and a Lys at the C-terminus. These were added to the crystal structure using Modloop (https://modbase.compbio.ucsf.edu/modloop/) and the 2 value dropped marginally to 9.2. Interestingly, in trying to fit the CaM data to the crystal structure, the FoXS calculation for CaM applies quite high values to the hydration layer scattering density, taking c2 to its limit of 4, which corresponds to ~0.388 electrons/Å3, in its attempt to minimize 2 for the unmodified crystal structure, and using a smaller but still high value for the modified structure (2.94). This value also is quite large for BSA (2.39), and smallest for GI (0.81). These values thus provide warning that the calculation is trying to adjust the hydration layer parameters for something missing in the model, which in these cases is likely flexibility based on complementary data from crystallography and NMR. Table S2 summarizes the results obtained using CRYSOL (Svergun et al., 1995) to optimize fits to the crystal structures, and as for FoXS, there is considerable variability in the applied adjustable parameters and the program seeks to minimize 2 for CaM using adjustable parameters relating to solvent subtraction and hydration layer thickness and density.

The overall mis-fits to the crystal structures for CaM and BSA is more clear in the error weighted residual difference plots than in the logI(q) vs q plot of the model overlaid with the experimental data (Fig. 2). Both BSA and CaM are two-domain structures, and the “wave” observed in the difference plot is suggestive of a relative shift, on average, in the relative positions of domains in solution compared to the crystal from.

The crystal structure of BSA shows two domains stabilized by a tight network of disulfides linked by a long flexible loop with high temperature factors assigned to residues 183–187 and 381–384 that are proposed to be responsible for domain movements (Pietrzyk et al., 2012). Allowing for flexibility in these residues and multi-state modelling with Multi-FoXS yielded a 2 1.05 for a 1-state model and the minimum 2 of 0.82 for a 3-state model. The model I(q) profiles for the 1- and 3- state models (Fig. 3A) fit within the noise, and even the residual difference plots between experiment and model I(q) are quite flat, with a possible narrowing of the difference plot for the 3-state model on the vertical scale (cyan symbols against black). Representative models from the best fit 1-, 2-, and 3-state models are shown in Fig. 3B, with the bead model from DAMMIN overlaid with the 1-state model and the crystal structure.

The alternate ensemble modelling program for flexible systems (EOM (Bernado et al., 2007)) was used model CaM accounting for the missing N- and C-terminal residues in and the known flexibility in the extended helix that connects the two globular domains of CaM (from NMR relaxation (Barbato et al., 1992) and solution SAXS (Heidorn and Trewhella, 1988) studies). A significantly lower 2 (0.82) is obtained and the model I(q) profile (Fig. 4A) fits within the noise, and as for the multi-state fits for BSA,  the residual difference plots between experiment and model I(q) are quite flat. From EOM, a quite large number of representative structures are selected (Fig. 4B) and this structural diversity is reflected in the frequency distributions for Rg and dmax (Figs 4D and C) in the optimized ensemble.

The modelling supports the conclusion from the Kratky plots that indicated BSA and CaM are each a mostly folded protein with some flexibility and is consistent with what is known about the proteins from crystallography and NMR. Of note, the P-values obtained from the CorMap analysis (Franke et al., 2015) support the ranking of goodness of fit based on 2. Interestingly, while 2 values are all relatively close to 1, and the error weighted difference plots were not strikingly helpful in demonstrating improvements in the fits beyond the 1-state models, the P-values do distinguish the models. The 1-state CaM model fit has a significantly lower P-value/adjusted P-value, with the longest run of consecutive points being on one side of the fit being 13 (locus at ~0.075 Å-1) versus 8 (locus ~0.03Å-1) for the 1-state versus 2-state models, respectively. For BSA the results are less clear. While P-values systematically increase from the 1- state to 2-state to 3-state models, the longest run of consecutive points is 14 (locus 0.2 Å-1), 12 (locus 0.01 Å-1), and 11 (locus 0.25 Å-1), respectively. These differences seem subtle and also, they occur in the lowest q and high q regimes, unlike the CaM example where the loci are more in the mid-q regime that we expect to be sensitive to domain movements.

Conclusions

The information provided as an example in Table 1, accompanied by Figs 1-3, provides a complete set of the information on the example SEC-SAXS experiments on GI, BSA and CaM in line with the recommended guidelines. The examples illustrate the value of comprehensive reporting for data and model in terms understanding data quality and model accuracy. Reporting of the often ignored adjustable parameters enhances understanding of potential limitations in the models presented. In this regard, it is noted that for some programs it is not straightforward to relate the adjustable parameters to the physical model. It would be desirable for developers of programs for SAS modelling to make this information more readily available. It is hoped that programmers might consider the reporting guidelines here and make readily available the necessary information and parameters in their standard output files.

Modern publishing practices may well require much of the reporting to be included as supplemental material. Eventually, most of it could be made available in the developing data and model archives. The latter will be increasingly important for managing large numbers of related data sets, although Figures 2 through 5 in (Carter et al., 2015) show how effectively one can assemble the results for multiple data sets in a form consistent with what is proposed here.

The three model proteins highlight advances in SEC-SAXS and the analysis of multi-state ensembles.  The aged BSA sample taken without any steps to remove higher oligomers prior to the SEC-SAXS measurement demonstrates power of SEC-SAXS. With careful attention to solvent subtraction, combined with M estimates from multiple methods confirming the size of the scattering particle was as predicted from chemical composition allowed the analysis to progress on this system with some confidence.  However,  it should be emphasized that starting with less than optimally prepared or poorly characterized samples for less well-known systems is unlikely to lead to results that can be interpreted with confidence. Where ever possible, high-resolution chromatography prior to SAXS or SEC-SAXS measurement is highly recommended.

As biomolecular SAS continues to grow and develop in this era of integrative/hybrid methods for structure determination of increasingly complex biomolecular complexes and assemblies, it is essential to firmly establish publication guidelines with the goal of ensuring access to the information required for proper evaluation of the quality of SAS samples and data, as well as the validity of structural interpretation. In addition to our recommended guidelines for data presentation in a publication, we recommend that SAS data be deposited and made freely available in a public data bank. Ideally q, I(q) with standard errors should be deposited for each measured profile, associated models plus the detail of how the experiment was conducted with the data and model validation parameters and analyses as outlined above and in the template Table S1. Our example data sets have been submitted to SASBDBD https://www.sasbdb.org/ that also archives relevant experimental conditions, sample details, derived models and their fits to the data.

In our examples, we have reported the P-values from the new CorMap analysis. At this stage, there is not broad experience to provide a basis for understanding the significance of the absolute value of the P-values in the context of SAS modelling. The continued accumulation of examples will build that necessary understanding and the SASBDB is collecting those examples. Once a sufficiently large sample size is acquired, there needs to be a systematic review so that a recommendation might be made regarding their use in reporting.

To support an envisioned future where the hybrid data supporting integrative/hybrid models with agreed validation protocols will be made available to the public (Sali et al., 2015), it also is recommended that SAS data banks utilize the sasCIF dictionary (Kachala et al., 2016) that enables seamless data exchange and interoperability with other structural biology data banks, including the wwPDB.

Author contributions:

All authors have reviewed, discussed and contributed to the preparation of the manuscript

JT, AD, AHK, NMK, TMR were responsible for SAS data acquisition and preliminary evaluation

JT, DS, DD, PV, AEW contributed to SAS data analysis and modelling

Table 1: SAS Results for Glucose Isomerase (GI), Bovine Serum Albumin (BSA) and Ca2+-Calmodulin (CaM)

| SAXS data collection parameters |  |  |  |
| --- | --- | --- | --- |
| Instrumentation with reference for data acquisition and processing details | Australian synchrotron SAXS-WAXS Beamline, http://www.synchrotron.org.au/aussyncbeamlines/saxswaxs |  |  |
| Wavelength | 1.0332 Å |  |  |
| Beam geometry | 250 x 130 m, 2.683 m camera length |  |  |
| q-measurement range (Å-1 or nm-1) | 0.00663 – 0.3104 Å-1 |  |  |
| Absolute scaling method | Comparison with scattering from pure H2O |  |  |
| Sample configuration | SEC-SAXS with sheath-flow cell (Kirby et al., 2016) |  |  |
| Exposure time | 1s x 24 |  |  |
| Sample temperature (oC) | 22ºC |  |  |
| Sample details |  |  |  |
|  | GI (tetramer) | BSA | CaM |
| Source (catalogue number) | Hampton Res. (HR7-100) | Sigma Aldrich (A3294) | Bacterially expressed (Michie et al., 2016) |
| Organism, UniProt sequence ID (residues in construct) | Streptomyces rubiginosus P24300 (2-388) | Bos Taurus P02769 (25-607) | Xenopus laevis P62155 (2-149) |
| Loading concentration (mg/mL), injection volume (L), mL/min for the SEC column (5 x 150 Superdex S200) | 6, 30, 0.45 | 25, 35, 0.45 | 20.2, 35, 0.45 |
| Average concentration in combined frames from A280 (mg/mL) | 0.18 | 0.565 | 0.96 |
| Extinction coefficient (A280, 0.1% solution) | 1.075 | 0.646 | 0.178 |
| Molecular mass from sequence (Da) | 172758 (tetramer) | 66430 (monomer) | 16842 (monomer) |
| Solvent details (solvent blanks taken from SEC flow through prior to elution of protein) | Buffer A1 | Buffer A1 | Buffer A1 |
| Partial specific volume from sequence (ϑ cm3 g-1) | 0.732 | 0.732 | 0.716 |
| Particle contrast from sequence and solvent constituents (Δρ, 1010 cm-2) | 2.839 | 2.863 | 3.093 |
| Guinier analysis |  |  |  |
| I(0) (cm-1) (normalised to unit C in mg/mL) | 0.206 ± 0.0002 | 0.0743 ± 0.00008 | 0.028 ± 0.00003 |
| Rg (Å) | 32.87 ± 0.13 | 28.33 ± 0.12 | 21.74 ± 0.06 |
| qRg max (qmin = 0.00663 Å-1) | 1.3 | 1.3 | 1.3 |
| Quality of fit (Pearson’s R2, qmin = 0.0063 Å-1 to qRg max = 1.3) | 0.999 | 0.999 | 0.999 |
| Molecular Weight from I(0) (ratio to expected value from sequence)b | 182480 (1.06) | 65120 (0.98) | 22003 (1.31) |
| P(r) analysis |  |  |  |
| I(0) (cm-1) (normalised to unit C in mg/mL) | 0.208 ± 0.0002 | 0.0752 ± 0.00006 | 0.0278 ± 0.00003 |
| Rg (Å) | 32.65E ± 0.04 | 28.32 ± 0.03 | 22.2 ± 0.06 |
| dmax (Å) | 92 | 87 | 72 |
| q range (Å-1) | 0.0066-0.2427 | 0.0066-0.2816 | 0.0074-0.3104 |
| 2 (total estimate from GNOM) | 0.929 (0.94) | 0.858 (0.96) | 0.855 (0.91) |
| Molecular Weight from I(0) (ratio to expected value from sequence) | 182480 (1.06) | 65800 (0.99) | 21734 (1.29) |
| Porod Volume (Å-3) (Mr estimate from (2/3)Vp) constant adjustment as % of background | 229000 (152667) ~3 | 101000 (67333) ~8 | 25200 (16800) ~30 |
| Model fitting results |  |  |  |
| Bead Modelling – DAMMIF |  |  |  |
| q range for fitting (Å-1) | 0.0066-0.2427 | 0.0066-0.2820 | 0.0074-0.3104 |
| Resolution (Å) (from DAMRES) | 37 ± 3 | 32 ± 3 | 30 ± 3 |
| Symmetry, anisotropy assumptions | P1, none | P1, none | P1, prolate |
| Ambiguity measure - Normalised Spatial Discrepancy, NSD (standard deviation) between models, number of clusters for 20 calculations for NSD > 0.7 | 0.62 (0.01) | 0.75 (0.63), 6 | 0.77 (0.02), 4 |
| 2 range for 20 calculations constant adjustment to intensities | 2.25 – 2.29 Skipped, unable to determine constant | 0.96 – 0.99 1.51e-4 | 1.30 – 1.37 1.48E-4 |
| M estimate for shape models (Da) | 134000 | 66700 | 16300 |
| Bead Modelling – DAMMIN |  |  |  |
| q range for fitting (Å-1) | 0.0066-0.2427 | 0.0066-0.2820 | 0.0074-0.3104 |
| Symmetry, anisotropy assumptions | P1 | P1 | P1 |
| 2, P /adj P (from CorMap) constant adjustment to intensities | 0.95, 0.035/0.722 2.697E-5 | 85 1.02E-4 | 0.844, 0.533/0.994 1.877E-4 |
| Atomistic Modelling- FoXS, MultiFoXS, EOM |  |  |  |
| 2, P /adj P (from CorMap), Predicted Rg (Å) adjustable parameters c1,c2 2 crystal structures: GI – PDB 1OAD (tetramer, two unique molecules from crystal structure, w. applied pseudo symm. dihedral D2) BSF – PDB 4F5S (monomer, Chain A) CaM – PDB1CLL and 1CLL+N-termADQ+C-term K (1CLL+) | 1.02, 0.0455/0.349 31.70 1.03, 0.81 | 4.4, 0.000/0.000 26.75 0.99, 2.39 | PDB:1CLL 1CLL: 10.8, 0.000/0.000 21.93 1.00, 4.00 |
|  |  |  | PDB 1CLL+ 9.12, 0.000/0.000 21.58 0.99, 2.94 |
| 2, P /adj P (from CorMap) for multi-state/ensemble modelling Multi-FoXS results for BSA – PDB:4F5S Chain A flexible residues 183-187 & 381-384. Adjustable parameters c1,c22 Weights, wn,, Rg (Å) for representative states 10,000 models for the initial ensemble. EOM results for CaM – PDB 1CLL + flexible residues 1-3 (ADQ), 77-87 (KDTDS). 10,000 models in the initial ensemble, native-like linkers selected, constant subtraction allowed. | n.a. | 1 state model 1.05, 0.023/0.129 0.99, 0.63 1 | 0.82, 0.788/0.998 constant subtraction = 0 13 representative structures chosen |
|  |  | 2 state model 0.96, 0.0891/0.5664 1.02, 1.21 0.833, 0.167 |  |
|  |  | 3 state model 0.82, 0.1709/0.9797 1.02, 0.94 0.74, 0.081, 0.079 |  |
| SASBDB ID for data and models | SASDCK2 https://www.sasbdb.org/data/SASDCK2/riiotttikh/ | to be submitted | SASDCQ23 https://www.sasbdb.org/data/SASDB25/29sa0zobpw/ |
| Software employed for SAXS data reduction, analysis and interpretation |  |  |  |
| SAXS data reduction to sample minus solvent, including desmearing, extrapolation, merging, etc as relevant |  | ScatterBrain 8.224 (http://www.synchrotron.org.au/aussyncbeamlines/saxswaxs/software-saxswaxs) for data reduction to I(q) vs q and PRIMUS from ATSAS 2.8.0 (Petoukhov et al., 2012) for solvent subtraction |  |
| Intensity scaling method |  | Absolute scaling based on H2O scattering |  |
| Extinction coefficient estimate |  | ProtParam (Gasteiger et al., 2005) |  |
| Calculation of Δρ and ϑ values from sequence |  | MULCh 1.1 (06/10/16) (Whitten et al., 2008) |  |
| Basic analyses: Guinier, P(r), Porod volume |  | PRIMUSQT from ATSAS 2.8.0 program suite (Petoukhov et al., 2012) |  |
| Shape/bead modelling |  | DAMMIF (Franke and Svergun, 2009) and DAMMIN (Svergun, 1999)via ATSAS on-line https://www.embl-hamburg.de/biosaxs/atsas-online/ |  |
| Atomic structure modelling (rigid body, ensemble) |  | FoXS (Schneidman-Duhovny et al., 2013) via web server (https://modbase.compbio.ucsf.edu/foxs/) CRYSOL 2.8.3 (Svergun et al., 1995) MultiFoXS (Schneidman-Duhovny et al., 2016) via web server https://modbase.compbio.ucsf.edu/multifoxs/ EOM (Bernado et al., 2007) via ATSAS on-line https://www.embl-hamburg.de/biosaxs/atsas-online/ |  |
| 3D graphic model representations |  | PyMOL v1.70.0.5 Win64 |  |

1 Buffer A: 25 mM MOPS, 250 mM NaCl, 50 mM KCl, 2 mM TCEP, 0.1% NaN3, pH 7.5

2 In FoXS the adjustable parameters c is the overall scaling factor in Eq. (5), c1 and c2 are adjustments for excluded volume and hydration density. c1 can vary by 5% (0.95-1.05) and the maximum hydration adjustment, normal density  = 0.334 electrons/Å3, while c2=4.0 corresponds to ~0.388 electrons/Å3.  In Multi-FoXs c1 and c2 are the same for all states, c is optimized for each state and a relative weight wn for each state n is output.

3 This submission was based on atomistic modelling for CaM using 1CLL without the additional N- and C-terminal residues.

4 This version of ScatterBrain outputs 2 x standard errors when manually processing SEC data frames, so the errors from ScatterBrain for the averaged measurement frames were halved to represent 1 x standard error propagation. Manual subtraction in PRIMUS preserves the 1 x standard error propagation.

Figure 1: SAS Results for three example proteins. Colour code: GI (blue), CaM (black) and BSA (red) A. Plot showing I(0) and Rg as a function of time for the SEC-SAXS run. B. I(q) vs q as a log-linear plot with the inset showing the Guinier fit for qRg <1.3. C. Dimensionless Kratky plot for the data in A. D. P(r) vs r from the data in A.

Figure 2: Crystal Structure Modelling Results for GI, BSA and CaM. FoXS-derived models (red and black solid lines) for PDB:1OAD, PDB:4F5S chain A and PDB:1CLL with the additional N- and C-terminal residues added fit to I(q) vs q for GI (blue), BSA (red) and CaM (black), respectively. The upper lot shows logI(q) vs q, while the lower inset is Δ =  Iexpq-cImod(q)σ(q)  vs q.

Figure 3: Multi-State Ensemble Modelling results for BSA. A. Model fits to I(q) vs q for BSA: I(q) vs q (black) for 1- (red) and 3- state (cyan) models assuming a flexible linker at residues 183-187 and 381-384. Error-weighted differences plots for the model and experiment profiles are shown above. The cyan dots in each case correspond to the best fit multi-state model (3-state for BSA), while the black corresponds to the 1-state model. B. Averaged and filtered DAMMIN model (wheat spheres) overlayed with the representative of the crystal structure (4F5S:Chain A, blue) and 1-state optimised model (magenta) for BSA, and representative structures from the 2- and 3- state optimised models (pink and cyan lines, respectively) for BSA.

Figure 3: Ensemble Optimized Modelling results for CaM  A. Model fits to I(q) vs q for CaM using PDB:1CLL and the full sequence: I(q) vs q (black) for 1- and 2- state (red line) assuming flexible residues 1-3 and 77-81. Error-weighted differences plot for the model and experiment profiles are shown above (red dots) B.  Averaged and filtered DAMMIN model (gray spheres) overlayed with representative structures from the optimized ensemble selected C. and D. Rg and dmax distributions, respectively, for the starting pool and optimized ensemble.

REFERENCES

Allen, A.J., Zhang, F., Kline, R.J., Guthrie, W.F., and Ilavsky, J. (2017). NIST Standard Reference Material 3600: Absolute Intensity Calibration Standard for Small-Angle X-ray Scattering. Journal of applied crystallography 50, 462-474.

Appolaire, A., Girard, E., Colombo, M., Dura, M.A., Moulin, M., Hartlein, M., Franzetti, B., and Gabel, F. (2014). Small-angle neutron scattering reveals the assembly mode and oligomeric architecture of TET, a large, dodecameric aminopeptidase. Acta crystallographica Section D, Biological crystallography 70, 2983-2993.

Babu, Y.S., Bugg, C.E., and Cook, W.J. (1988). Structure of calmodulin refined at 2.2 A resolution. Journal of molecular biology 204, 191-204.

Barbato, G., Ikura, M., Kay, L.E., Pastor, R.W., and Bax, A. (1992). Backbone dynamics of calmodulin studied by 15N relaxation using inverse detected two-dimensional NMR spectroscopy: the central helix is flexible. Biochemistry 31, 5269-5278.

Bergmann, A., Orthaber, D., Scherf, G., and Glatter, O. (2000). Improvement of SAXS measurements on Kratky slit systems by Göbel mirrors and imaging-plate detectors. J Appl Crystallogr 33, 869-875.

Berlin, K., Castaneda, C.A., Schneidman-Duhovny, D., Sali, A., Nava-Tudela, A., and Fushman, D. (2013). Recovering a representative conformational ensemble from underdetermined macromolecular structural data. Journal of the American Chemical Society 135, 16595-16609.

Berman, H.M., Bhat, T.N., Bourne, P.E., Feng, Z., Gilliland, G., Weissig, H., and Westbrook, J. (2000). The Protein Data Bank and the challenge of structural genomics. Nature structural biology 7 Suppl, 957-959.

Bernado, P., Mylonas, E., Petoukhov, M.V., Blackledge, M., and Svergun, D.I. (2007). Structural characterization of flexible proteins using small-angle X-ray scattering. Journal of the American Chemical Society 129, 5656-5664.

Bernado, P., Perez, Y., Svergun, D.I., and Pons, M. (2008). Structural characterization of the active and inactive states of Src kinase in solution by small-angle X-ray scattering. Journal of molecular biology 376, 492-505.

Berthaud, A., Manzi, J., Perez, J., and Mangenot, S. (2012). Modeling detergent organization around aquaporin-0 using small-angle X-ray scattering. Journal of the American Chemical Society 134, 10080-10088.

Blanchet, C.E., Spilotros, A., Schwemmer, F., Graewert, M.A., Kikhney, A., Jeffries, C.M., Franke, D., Mark, D., Zengerle, R., Cipriani, F., et al. (2015). Versatile sample environments and automation for biological solution X-ray scattering experiments at the P12 beamline (PETRA III, DESY). Journal of applied crystallography 48, 431-443.

Bras, W., Koizumi, S., and Terrill, N.J. (2014). Beyond simple small-angle X-ray scattering: developments in online complementary techniques and sample environments. IUCrJ 1, 478-491.

Brennich, M.E., Round, A.R., and Hutin, S. (2017). Online Size-exclusion and Ion-exchange Chromatography on a SAXS Beamline. Journal of Visualized Experiments.

Brookes, E., Perez, J., Cardinali, B., Profumo, A., Vachette, P., and Rocco, M. (2013). Fibrinogen species as resolved by HPLC-SAXS data processing within the UltraScan Solution Modeler (US-SOMO) enhanced SAS module. Journal of applied crystallography 46, 1823-1833.

Brookes, E., Vachette, P., Rocco, M., and Perez, J. (2016). US-SOMO HPLC-SAXS module: dealing with capillary fouling and extraction of pure component patterns from poorly resolved SEC-SAXS data. Journal of applied crystallography 49, 1827-1841.

Carter, L., Kim, S.J., Schneidman-Duhovny, D., Stohr, J., Poncet-Montange, G., Weiss, T.M., Tsuruta, H., Prusiner, S.B., and Sali, A. (2015). Prion Protein-Antibody Complexes Characterized by Chromatography-Coupled Small-Angle X-Ray Scattering. Biophysical journal 109, 793-805.

Chen, P.C., and Hub, J.S. (2015). Interpretation of solution x-ray scattering by explicit-solvent molecular dynamics. Biophysical journal 108, 2573-2584.

David, G., and Perez, J. (2009). Combined sampler robot and high-performance liquid chromatography: A fully automated system for biological small-angle X-ray scattering experiments at the Synchrotron SOLEIL SWING beamline. J Appl Crystallogr 42, 892-900.

Debye, P. (1947). Molecular weight determined by light scattering. J Phys Colloid Chem 51, 18-32.

Debye, P., Andrerson, H.R., and Brumberger, H. (1957). Scattering by an inhomogeneous solid 2. The correlation function and its application. J Appl Phys 20, 679-683.

Fiegen, L.A., and Svergun, D.I. (1987). Structure Analysis by Small-Angle X-ray and Neutron Scattering (New York: Plenum Press).

Fischer, H., de Oliveira Neto, M., B., N.H., Polikarpov, I., and Craievich, A.F. (2009). J Appl Crystallogr 43, 101–109.

Franke, D., Jeffries, C.M., and Svergun, D.I. (2015). Correlation Map, a goodness-of-fit test for one-dimensional X-ray scattering spectra. Nat Methods 12, 419-422.

Franke, D., and Svergun, D.I. (2009). DAMMIF, a program for rapid ab-initio shape determination in small-angle scattering. Journal of applied crystallography 42, 342-346.

Gabel, F. (2012). A simple procedure to evaluate the efficiency of bio-macromolecular rigid-body refinement by small-angle scattering. European biophysics journal : EBJ 41, 1-11.

Gasteiger, E., Hoogland, C., Gattiker, A., Duvaud, S., Wilkins, M.R., Appel, R.D., and Bairoch, A. (2005). Protein Identification and Analysis Tools on the ExPASy Server. In The Proteomics Protocols Handbook, J.M. Walker, ed. (New York: Humana Press), pp. 571-607.

Glatter, O. (1977). A new method for the evaluation of small-angle scattering data. J Appl Crystallogr 10, 307-315.

Glatter, O., and Kratky, O. (1982). Small-Angle X-ray Scattering (London: Academic Press).

Graewert, M.A., Franke, D., Jeffries, C.M., Blanchet, C.E., Ruskule, D., Kuhle, K., Flieger, A., Schafer, B., Tartsch, B., Meijers, R., et al. (2015). Automated pipeline for purification, biophysical and x-ray analysis of biomacromolecular solutions. Scientific reports 5, 10734.

Grishaev, A., Guo, L., Irving, T., and Bax, A. (2010). Improved fitting of solution X-ray scattering data to macromolecular structures and structural ensembles by explicit water modeling. Journal of the American Chemical Society 132, 15484-15486.

Guinier, A. (1939). La diffraction des rayons x aux très faibles angles: Applications à l'etude de phénomènes ultra-microscopiques. Ann Phys-Paris 12, 161-237.

Heidorn, D.B., and Trewhella, J. (1988). Comparison of the crystal and solution structures of calmodulin and troponin C. Biochemistry 27, 909-915.

Ibrahim, Z., Martel, A., Moulin, M., Kim, H.S., Hartlein, M., Franzetti, B., and Gabel, F. (2017). Time-resolved neutron scattering provides new insight into protein substrate processing by a AAA+ unfoldase. Scientific reports 7, 40948.

Jacques, D.A., Guss, J.M., Svergun, D.I., and Trewhella, J. (2012a). Publication guidelines for structural modelling of small-angle scattering data from biomolecules in solution. Acta crystallographica Section D, Biological crystallography 68, 620-626.

Jacques, D.A., Guss, J.M., and Trewhella, J. (2012b). Reliable structural interpretation of small-angle scattering data from bio-molecules in solution--the importance of quality control and a standard reporting framework. BMC Struct Biol 12, 9.

Jacrot, B. (1976). The Study of Biological Structures in Neutron Scattering from Solution. Reports on Progress in Physics 39, 911-953.

Jacrot, B., and Zaccai, G. (1981). Determination of molecular weight by neutron scattering. Biopolymers 20, 2413–2426.

Jeffries, C.M., Graewert, M.A., Blanchet, C.E., Langley, D.B., Whitten, A.E., and Svergun, D.I. (2016). Preparing monodisperse macromolecular samples for successful biological small-angle X-ray and neutron-scattering experiments. Nature protocols 11, 2122-2153.

Jordan, A., Jacques, M., Merrick, C., Devos, J., Forsyth, V.T., Porcar, L., and Martel, A. (2016). SEC-SANS: size exclusion chromatography combined in situ with small-angle neutron scattering. Journal of applied crystallography 49, 2015-2020.

Kachala, M., Westbrook, J., and Svergun, D. (2016). Extension of the sasCIF format and its applications for data processing and deposition. Journal of applied crystallography 49, 302-310.

Kikhney, A.G., and Svergun, D.I. (2015). A practical guide to small angle X-ray scattering (SAXS) of flexible and intrinsically disordered proteins. FEBS letters 589, 2570-2577.

Kim, H.S., and Gabel, F. (2015). Uniqueness of models from small-angle scattering data: the impact of a hydration shell and complementary NMR restraints. Acta crystallographica Section D, Biological crystallography 71, 57-66.

Kirby, N., Cowieson, N., Hawley, A.M., Mudie, S.T., McGillivray, D.J., Kusel, M., Samardzic-Boban, V., and Ryan, T.M. (2016). Improved radiation dose efficiency in solution SAXS using a sheath flow sample environment. Acta crystallographica Section D, Structural biology 72, 1254-1266.

Koch, M.H., and Stuhrmann, H.B. (1979). Neutron-scattering studies of ribosomes. Methods in enzymology 59, 670-706.

Konarev, P.V., and Svergun, D.I. (2015). A posteriori determination of the useful data range for small-angle scattering experiments on dilute monodisperse systems. IUCrJ 2, 352-360.

Konarev, P.V., Volkov, V.V., Sokolova, A.V., Koch, M.H.J., and Svergun, D.I. (2003). PRIMUS: a Windows PC-based system for small-angle scattering data analysis. J Appl Cryst 36, 1277-1282.

Kratky, O. (1982). Natural High Polymers in the Dissolved and Solid State. In Small-Angle X-ray Scattering, O. Glatter, and O. Kratky, eds. (London: Academic Press), pp. 361-386.

Malfois, M., and Svergun, D.I. (2000). sasCIF: an extension of core Crystallographic Information File for SAS. J Appl Cryst 33, 812–816.

Mathew, E., Mirza, A., and Menhart, N. (2004). Liquid-chromatography-coupled SAXS for accurate sizing of aggregating proteins. J Synchrotron Radiat 11, 314-318.

Meisburger, S.P., Taylor, A.B., Khan, C.A., Zhang, S., Fitzpatrick, P.F., and Ando, N. (2016). Domain Movements upon Activation of Phenylalanine Hydroxylase Characterized by Crystallography and Chromatography-Coupled Small-Angle X-ray Scattering. Journal of the American Chemical Society 138, 6506-6516.

Michie, K.A., Kwan, A.H., Tung, C.S., Guss, J.M., and Trewhella, J. (2016). A Highly Conserved Yet Flexible Linker Is Part of a Polymorphic Protein-Binding Domain in Myosin-Binding Protein C. Structure 24, 2000-2007.

Mylonas, E., and Svergun, D.I. (2007). Accuracy of molecular mass determination of proteins in solution by small-angle X-ray scattering. J Appl Cryst 40, s245–s249.

Orthaber, D., Bergmann, A., and Glatter, O. (2000). SAXS experiments on absolute scale with Kratky systems using water as a secondary standard. J Appl Cryst 33, 218-225.

Pelikan, M., Hura, G.L., and Hammel, M. (2009). Structure and flexibility within proteins as identified through small angle X-ray scattering. General Physiology and Biophysics 28, 174–189.

Perkins, S.J. (1986). Protein volumes and hydration effects. The calculations of partial specific volumes, neutron scattering matchpoints and 280-nm absorption coefficients for proteins and glycoproteins from amino acid sequences. European journal of biochemistry 157, 169-180.

Perkins, S.J., Wright, D.W., Zhang, H., Brookes, E.H., Chen, J., Irving, T.C., Krueger, S., Barlow, D.J., Edler, K.J., Scott, D.J., et al. (2016). Atomistic modelling of scattering data in the Collaborative Computational Project for Small Angle Scattering (CCP-SAS). Journal of applied crystallography 49, 1861-1875.

Petoukhov, M.V., Billas, I.M., Takacs, M., Graewert, M.A., Moras, D., and Svergun, D.I. (2013). Reconstruction of quaternary structure from X-ray scattering by equilibrium mixtures of biological macromolecules. Biochemistry 52, 6844-6855.

Petoukhov, M.V., Franke, D., Shkumatov, A.V., Tria, G., Kikhney, A.G., Gajda, M., Gorba, C., Mertens, H.D., Konarev, P.V., and Svergun, D.I. (2012). New developments in the ATSAS program package for small-angle scattering data analysis. Journal of applied crystallography 45, 342-350.

Petoukhov, M.V., and Svergun, D.I. (2015). Ambiguity assessment of small-angle scattering curves from monodisperse systems. Acta crystallographica Section D, Biological crystallography 71, 1051-1058.

Pietrzyk, A.J., Panjikar, S., Bujacz, A., Mueller-Dieckmann, J., Lochynska, M., Jaskolski, M., and Bujacz, G. (2012). High-resolution structure of Bombyx mori lipoprotein 7: crystallographic determination of the identity of the protein and its potential role in detoxification. Acta crystallographica Section D, Biological crystallography 68, 1140-1151.

Porod, G. (1951). Die Röntgenkleinwinkelstreung von dichtgepackten kolloidalen Systemen. Kolloid Z 124, 83-114.

Ramagopal, U.A., Dauter, M., and Dauter, Z. (2003). SAD manganese in two crystal forms of glucose isomerase. Acta crystallographica Section D, Biological crystallography 59, 868-875.

Rambo, R.P., and Tainer, J.A. (2011). Characterizing flexible and intrinsically unstructured biological macromolecules by SAS using the Porod-Debye law. Biopolymers 95, 559-571.

Rambo, R.P., and Tainer, J.A. (2013). Accurate assessment of mass, models and resolution by small-angle scattering. Nature 496, 477-481.

Sali, A., Berman, H.M., Schwede, T., Trewhella, J., Kleywegt, G., Burley, S.K., Markley, J., Nakamura, H., Adams, P., Bonvin, A.M., et al. (2015). Outcome of the First wwPDB Hybrid/Integrative Methods Task Force Workshop. Structure 23, 1156-1167.

Schneidman-Duhovny, D., Hammel, M., Tainer, J.A., and Sali, A. (2013). Accurate SAXS profile computation and its assessment by contrast variation experiments. Biophysical journal 105, 962-974.

Schneidman-Duhovny, D., Hammel, M., Tainer, J.A., and Sali, A. (2016). FoXS, FoXSDock and MultiFoXS: Single-state and multi-state structural modeling of proteins and their complexes based on SAXS profiles. Nucleic acids research 44, W424-429.

Schneidman-Duhovny, D., Pellarin, R., and Sali, A. (2014). Uncertainty in integrative structural modeling. Current opinion in structural biology 28, 96-104.

Schwieters, C.D., and Clore, G.M. (2014). Using small angle solution scattering data in Xplor-NIH structure calculations. Prog Nucl Magn Reson Spectrosc 80, 1-11.

Sedlak, S.M., Bruetzel, L.K., and Lipfert, J. (2017). Quantitative evaluation of statistical errors in small-angle X-ray scattering measurements. Journal of applied crystallography 50, 621-630.

Skou, S., Gillilan, R.E., and Ando, N. (2014). Synchrotron-based small-angle X-ray scattering of proteins in solution. Nature protocols 9, 1727-1739.

Stuhrmann, H.B. (1980). Small-Angle X-Ray Scattering of Macromolecules in Solution. In Synchrotron Radiation Research, H. Winick, and S. Doniach, eds. (New York: Springer US), pp. 513-531.

Svergun, D., Barberato, C., and Koch, M.H.J. (1995). CRYSOL - A program to evaluate x-ray solution scattering of biological macromolecules from atomic coordinates. Journal of applied crystallography 28, 768-773.

Svergun, D.I. (1992). Determination of the regularization parameter in indirect-transform methods using perceptual criteria. J Appl Crystallogr 25, 495–503.

Svergun, D.I. (1999). Restoring low resolution structure of biological macromolecules from solution scattering using simulated annealing. Biophysical journal 76, 2879-2886.

Svergun, D.I., Koch, M.H.J., Timmins, P.A., and May, R.P. (2013). Small-Angle X-ray and Neutron Scattering from Biologiccal Macromolecules, Vol 19 (Oxford University Press).

Svergun, D.I., Richard, S., Koch, M.H., Sayers, Z., Kuprin, S., and Zaccai, G. (1998). Protein hydration in solution: experimental observation by x-ray and neutron scattering. Proceedings of the National Academy of Sciences of the United States of America 95, 2267-2272.

Terakawa, T., Higo, J., and Takada, S. (2014). Multi-scale ensemble modeling of modular proteins with intrinsically disordered linker regions: application to p53. Biophysical journal 107, 721-729.

Trewhella, J. (2016). Small-angle scattering and 3D structure interpretation. Current opinion in structural biology 40, 1-7.

Trewhella, J., Hendrickson, W.A., Kleywegt, G.J., Sali, A., Sato, M., Schwede, T., Svergun, D.I., Tainer, J.A., Westbrook, J., and Berman, H.M. (2013). Report of the wwPDB Small-Angle Scattering Task Force: data requirements for biomolecular modeling and the PDB. Structure 21, 875-881.

Tria, G., Mertens, H.D., Kachala, M., and Svergun, D.I. (2015). Advanced ensemble modelling of flexible macromolecules using X-ray solution scattering. IUCrJ 2, 207-217.

Tuukkanen, A.T., Kleywegt, G.J., and Svergun, D.I. (2016). Resolution of ab initio shapes determined from small-angle scattering. IUCrJ 3, 440-447.

Valentini, E., Kikhney, A.G., Previtali, G., Jeffries, C.M., and Svergun, D.I. (2015). SASBDB, a repository for biological small-angle scattering data. Nucleic acids research 43, D357-363.

Whitten, A.E., Cai, S.Z., and Trewhella, J. (2008). MULCh: modules for the analysis of small-angle neutron contrast variation data from biomolecular assemblies. J Appl Crystallogr 41, 222-226.

Yang, S., Blachowicz, L., Makowski, L., and Roux, B. (2010). Multidomain assembled states of Hck tyrosine kinase in solution. Proceedings of the National Academy of Sciences of the United States of America 107, 15757-15762.

Yang, Z., Lasker, K., Schneidman-Duhovny, D., Webb, B., Huang, C.C., Pettersen, E.F., Goddard, T.D., Meng, E.C., Sali, A., and Ferrin, T.E. (2012). UCSF Chimera, MODELLER, and IMP: an integrated modeling system. Journal of structural biology 179, 269-278.

Zhang, F., Roosen-Runge, F., Skoda, M.W., Jacobs, R.M., Wolf, M., Callow, P., Frielinghaus, H., Pipich, V., Prevost, S., and Schreiber, F. (2012). Hydration and interactions in protein solutions containing concentrated electrolytes studied by small-angle scattering. Physical chemistry chemical physics : PCCP 14, 2483-2493.